## Pin-map_HIFU_V1.0
Rubbing HIFU Block Diagram V1.0
|  | CN102 |  | HD1 |  |  |  | HD1 |  |  |  | HD3 |  |  |  | HD1 |  |  |  |  | RF_OUT1
|  | 5267-2P |  | RF_QR_IN | 1 | PW_IS_CT | 2 | 1 | RF_QR_IN | 2 | PW_IS_CT | GND | 1 | GND | 2 | 1 | GND | 2 | GND |  | SMA-R-R/A(BNC)
|  | 1 | FOOT_SW | BIAS_ON | 3 | VDC_OUT | 4 | 3 | BIAS_ON | 4 | VDC_OUT | GND | 3 | GND | 4 | 3 | GND | 4 | GND |  | HIFU_BODY | 1
|  | 2 | GND | TIM8_OC_IN | 5 | VDC_IN | 6 | 5 | TIM8_OC_IN | 6 | VDC_IN | - | 5 | - | 6 | 5 | - | 6 | - |  | GND | 2
|  |  |  | PW_PWM | 7 | FW_CT | 8 | 7 | PW_PWM | 8 | FW_CT | - | 7 | - | 8 | 7 | - | 8 | -
|  | FAN1 |  | VALVE_ON | 9 | FET_TEMP | 10 | 9 | VALVE_ON | 10 | FET_TEMP | - | 9 | - | 10 | 9 | - | 10 | - |  | RF_OUT2
|  | 5267-2P |  | C_FAN_ON | 11 | - | 12 | 11 | C_FAN_ON | 12 | - | - | 11 | - | 12 | 11 | - | 12 | - |  | SMA-R-R/A(BNC)
|  | 1 | +12V | PUMP_ON | 13 | PW_I_REF | 14 | 13 | PUMP_ON | 14 | PW_I_REF | - | 13 | - | 14 | 13 | - | 14 | - |  | HIFU_FACE | 1
|  | 2 | FAN_PWM | - | 15 | - | 16 | 15 | - | 16 | - | - | 15 | - | 16 | 15 | - | 16 | - |  | GND | 2
|  |  |  | RELAY_ON | 17 | FT_7M_RY_ON | 18 | 17 | RELAY_ON | 18 | FT_7M_RY_ON | - | 17 | - | 18 | 17 | - | 18 | -
|  | FAN2 |  | RF_A_RY_ON | 19 | FT_4M_RY_ON | 20 | 19 | RF_A_RY_ON | 20 | FT_4M_RY_ON | RF_OUT | 19 | RF_OUT | 20 | 19 | RF_Q_IN | 20 | RF_Q_IN
|  | 5267-2P |  | RF_B_RY_ON | 21 | FT_2M_RY_ON | 22 | 21 | RF_B_RY_ON | 22 | FT_2M_RY_ON
|  | 1 | +12V | - | 23 | GND | 24 | 23 | - | 24 | GND | HD4 |  |  |  | HD2
|  | 2 | FAN_PWM | - | 25 | GND | 26 | 25 | - | 26 | GND | RY_HL2_7M | 1 | RY_HL2_7M | 2 | 1 | RY_HL2_7M | 2 | RY_HL2_7M
[12 empty] | RY_HL2_4M | 3 | RY_HL2_4M | 4 | 3 | RY_HL2_4M | 4 | RY_HL2_4M
|  | CN101 | [9 empty] | RY_HL2_2M | 5 | RY_HL2_2M | 6 | 5 | RY_HL2_2M | 6 | RY_HL2_2M
|  | 5267-3P |  | HP1 | [7 empty] | RY_HL2_A | 7 | RY_HL2_A | 8 | 7 | RY_HL2_A | 8 | RY_HL2_A
|  | 1 | TXD_LCD | SHOT_SW_A | 1 | HP_OK_A | 2 |  |  |  |  | RY_HL2_B | 9 | RY_HL2_B | 10 | 9 | RY_HL2_B | 10 | RY_HL2_B
|  | 2 | RXD_LCD | KART_CHK_A | 3 | - | 4 |  |  |  |  | - | 11 | - | 12 | 11 | - | 12 | -
|  | 3 | GND | HALL_SW_A | 5 | AA5V | 6 |  |  |  |  | GND | 13 | GND | 14 | 13 | GND | 14 | GND
|  |  |  | TXD_ACT_A | 7 | CABLE_SDA_A | 8 |  |  |  |  | GND | 15 | GND | 16 | 15 | GND | 16 | GND
|  | CN106 |  | RXD_ACT_A | 9 | CABLE_SCL_A | 10 |  |  |  |  | - | 17 | - | 18 | 17 | - | 18 | -
|  | MX35313-2P |  | AA12V | 11 | GND | 12 |  |  |  |  | FW_CT | 19 | FW_CT | 20 | 19 | FW_CT | 20 | FW_CT
|  | 1 | +12V | AA12V | 13 | GND | 14
|  | 2 | GND
HP2
|  | JTAG2 |  | SHOT_SW_B | 1 | HP_OK_B | 2 |  |  |  |  | HD2 |  |  |  | HD1
|  | YDW200-06P |  | KART_CHK_B | 3 | - | 4 |  |  |  |  | SS48V | 1 | SS48V | 2 | 1 | +48V | 2 | +48V
|  | 1 | 3_3V | HALL_SW_B | 5 | BB5V | 6 |  |  |  |  | SS48V | 3 | SS48V | 4 | 3 | +48V | 4 | +48V
|  | 2 | SCI_RXA | TXD_ACT_B | 7 | CABLE_SDA_B | 8 |  |  |  |  | - | 5 | - | 6 | 5 | - | 6 | -
|  | 3 | BOOT0 | RXD_ACT_B | 9 | CABLE_SCL_B | 10 |  |  |  |  | - | 7 | - | 8 | 7 | - | 8 | -
|  | 4 | SCI_TXA | BB12V | 11 | GND | 12 |  |  |  |  | GND | 9 | GND | 10 | 9 | GND | 10 | GND
|  | 5 | GND | BB12V | 13 | GND | 14 |  |  |  |  | GND | 11 | GND | 12 | 11 | GND | 12 | GND |  | CN202
|  | 6 | NRST | [8 empty] | - | 13 | - | 14 | 13 | +12V | 14 | +12V |  | MOLEX53259-2P
[12 empty] | PUMP_DR | 15 | PUMP_DR | 16 | 15 | PUMP_DR | 16 | PUMP_DR |  | +12V | 1
[12 empty] | CHILER_FAN_DR | 17 | CHILER_FAN_DR | 18 | 17 | CHILER_FAN_DR | 18 | CHILER_FAN_DR |  | VALVE_DR | 2
[12 empty] | VALVE_DR | 19 | VALVE_DR | 20 | 19 | VALVE_DR | 20 | VALVE_DR
[12 empty] | - | 21 | - | 22 | 21 | - | 22 | - |  | CN203
[12 empty] | RY7 input | 23 | RY7 input | 24 | 23 | Pin-24 | 24 | Pin-23 |  | LWB1143-2P
[12 empty] | RY7 output | 25 | RY7 output | 26 | 25 | Pin-26 | 26 | Pin-25 |  | +12V | 1
[21 empty] | PUMP_DR | 2
[6 empty] | CN103 | [9 empty] | CN204
[6 empty] | 5267-3P | [9 empty] | MOLEX5267-2P
[6 empty] | F_5V | 1 | [8 empty] | 1 | +12V
[6 empty] | GND | 2 | [8 empty] | 2 | CHILER_FAN_DR
[6 empty] | FLOW_SENSOR | 3
CN205
[6 empty] | CN104 | [9 empty] | MOLEX5267-2P
CN401 | [5 empty] | 5267-3P | [9 empty] | 1 | +12V
5267-5P | [5 empty] | WATER_LV1 | 1 | [8 empty] | 2 | CHILER_FAN_DR
1 | +5V |  |  |  |  | WATER_LV2 | 2
2 | BLUE_LED |  |  |  |  | GND | 3
3 | RED_LED | [10 empty] | CN403 |  |  |  | CN201
4 | [5 empty] | CN105 |  | CN402 |  |  |  | MX5566-04P |  |  |  | MX5569-04P
5 | GND |  |  |  |  | MX35313-2P |  | LAB1143-2P |  |  |  | S12V | 1 | GND | 2 | 1 | +12V | 2 | GND
[6 empty] | +12V | 1 | 1 | S12V |  |  | S24V | 3 | GND | 4 | 3 | +48V | 4 | GND
CN401 | [5 empty] | GND | 2 | 2 | GND
1 | +5V | [12 empty] | CN405 | [5 empty] | CN2
2 | BLUE_LED | [12 empty] | 35313-2P | [5 empty] | MOLEX-35313-3P
3 | RED_LED |  |  |  |  | CN401 |  | CN401 | [5 empty] | +12V | 1 |  |  |  |  | 1 | L
4 | [5 empty] | 5267-5P |  | 5267-5P | [5 empty] | GND | 2 |  |  |  |  | 2
5 | GND |  |  |  |  | +5V | 1 | 1 | +5V | [10 empty] | 3 | N |  | CN1
[6 empty] | BLUE_LED | 2 | 2 | BLUE_LED | [13 empty] | MOLEX-35313-5P
[6 empty] | RED_LED | 3 | 3 | RED_LED |  |  |  |  | CN404 | [5 empty] | CN3 |  |  | L | 1
[6 empty] | S/W | 4 | 4 | S/W |  |  |  |  | 35313-3P | [5 empty] | MOLEX-35313-3P |  |  |  | 2
[6 empty] | GND | 5 | 5 | GND |  |  |  |  | +24V | 1 |  |  |  |  | 1 | L |  | N | 3
[14 empty] | - | 2 |  |  |  |  | 2 |  |  |  | 4
[14 empty] | GND | 3 |  |  |  |  | 3 | N |  | FG | 5
VDM211
(DC12V)
15RNS-L
(DC12V)
H/P
(BODY)
H/P
(FACE)
PWR
CPU
H/P
(Body)
H/P
(Face)
FOOT
S/W
FAN 1
(PWR)
FAN 2
(PWR)
LCD
10.2 inch
(ForStech)
FRONT LED
SMPS 12V
(MDS-100BPS12)
SMPS 48V
(MDS-300APB48)
CT Block
KEY LED
FLT
FLOW SENSOR
()
WATER TANK
(water level sensor)
PCN
FAN 4
(CHILER)
FAN 3
(CHILER)
DC/DC
(LM2576HV-12)
DC 12V→9V
(7809AF)
+9V
+12V
SS48V
+12V
CHILER
LSW
MagicPolar LF
LED

## Pin-map_Focus
Focus Dual Block Diagram
| CN101 |  | HD3 |  |  |  | HD1 |  |  |  | HD2 |  |  |  |  | HD2 |  |  |  | RF_OUT |  | RF_IN |  | HD1 |  |  |  | HD1
| 5267-2P |  | RF_QR_IN | 1 | PW_IS_CT | 2 | 1 | RF_QR_IN | 2 | PW_IS_CT | RF_OUT | 1 | RF_OUT | 2 |  | 1 | RF_OUT | 2 | RF_OUT | SMA-R-R/A(BNC) |  | SMA-R-R/A(BNC) |  | RF_RY1 | 1 | FIL_RY_4M | 2 | 1 | RF_RY1 | 2
| 1 | FOOT_SW | PW_PWM | 3 | VDC_OUT | 4 | 3 | PW_PWM | 4 | VDC_OUT | - | 3 | - | 4 |  | 3 | - | 4 | - | RF_OUT | 1 | 1 | RF_IN | RF_RY2 | 3 | FIL_RY_7M | 4 | 3 | RF_RY2 | 4 | FIL_RY_7M
| 2 | GND | MOT_PWM | 5 | VDC_IN | 6 | 5 | MOT_PWM | 6 | VDC_IN | - | 5 | - | 6 |  | 5 | - | 6 | - | GND | 2 | 2 | GND | GND | 5 | GND | 6 | 5 | GND | 6 | GND
|  |  | PUMP_IS_CT | 7 | FW_CT | 8 | 7 | PUMP_IS_CT | 8 | FW_CT | GND | 7 | GND | 8 |  | 7 | GND | 8 | GND |  |  |  |  | +12V | 7 | +12V | 8 | 7 | +12V | 8 | +12V
| FAN1 |  | TIM1_OC_IN | 9 | RV_CT | 10 | 9 | - | 10 | RV_CT | GND | 9 | GND | 10 |  | 9 | GND | 10 | GND |  |  |  |  | - | 9 | - | 10 | 9 | - | 10 | -
| 5267-2P |  | TIM8_OC_IN | 11 | FET_TEMP | 12 | 11 | TIM8_OC_IN | 12 | FET_TEMP | - | 11 | - | 12 |  | 11 | - | 12 | - | CN202 |  |  |  | - | 11 | - | 12 | 11 | - | 12 | -
| 1 | +12V | BIAS_ON | 13 | PW_I_REF | 14 | 13 | BIAS_ON | 14 | PW_I_REF | - | 13 | - | 14 |  | 13 | - | 14 | - | 53259-2P |  |  |  | RF_IN | 13 | RF_IN | 14 | 13 | RF_IN | 14 | RF_IN
| 2 | FAN_ON | VALVE2_ON | 15 | VALVE1_ON | 16 | 15 | VALVE2_ON | 16 | VALVE1_ON | SS48V | 15 | SS48V | 16 | 오타 | 15 | +24V | 16 | +24V | +12V | 1 |  |  | GND | 15 | GND | 16 | 15 | GND | 16 | GND
|  |  | - | 17 | FILTER_7M | 18 | 17 | - | 18 | FILTER_7M | SS48V | 17 | SS48V | 18 |  | 17 | +24V | 18 | +24V | VAL1_COIL | 2 |  |  | RF_OUT1 | 17 | RF_OUT1 | 18 | 17 | RF_OUT1 | 18 | RF_OUT1
| FAN2 |  | HF_OUT1 | 19 | FILTER_4M | 20 | 19 | HF_OUT1 | 20 | FILTER_4M | - | 19 | - | 20 |  | 19 | - | 20 | - |  |  |  |  | GND | 19 | GND | 20 | 19 | GND | 20 | GND
| 5267-2P |  | HF_OUT2 | 21 | FILTER_2M | 22 | 21 | HF_OUT2 | 22 | FILTER_2M | [9 empty] | CN203 |  |  |  | - | 21 | - | 22 | 21 | - | 22 | -
| 1 | +12V | - | 23 | GND | 24 | 23 | - | 24 | GND | HD3 |  |  |  |  | HD1 |  |  |  | 53259-2P |  |  |  | - | 23 | - | 24 | 23 | - | 24 | -
| 2 | FAN_ON | - | 25 | GND | 26 | 25 | - | 26 | GND | VAL1_COIL | 1 | VAL1_COIL | 2 |  | 1 | VAL1_COIL | 2 | VAL1_COIL | +12V | 1 |  |  | - | 25 | - | 26 | 25 | - | 26 | -
[11 empty] | VAL2_COIL | 3 | VAL2_COIL | 4 |  | 3 | VAL2_COIL | 4 | VAL2_COIL | VAL2_COIL | 2 |  |  | NED_RF_A | 27 | NED_RF_B | 28 | 27 | NED_RF_A | 28 | NED_RF_B
| CN102 | [9 empty] | - | 5 | - | 6 |  | 5 | - | 6 | - |  |  |  |  | NED_RF_A | 29 | NED_RF_B | 30 | 29 | NED_RF_A | 30 | NED_RF_B
| 5267-3P |  | HD1 | [7 empty] | IS_MOT | 7 | IS_MOT | 8 |  | 7 | IS_MOT | 8 | IS_MOT | CN204
| 1 | TXD_LCD | HIF_SHOT_SW | 1 | HIF_HP_SW | 2 |  |  |  |  | GND | 9 | GND | 10 |  | 9 | GND | 10 | GND | LWB1143-2P | [5 empty] | RF_OUT1
| 2 | RXD_LCD | HIF_CART_SW | 3 | - | 4 |  |  |  |  | GND | 11 | GND | 12 |  | 11 | GND | 12 | GND | +12V | 1 |  |  |  |  | SMA-R-R/A(BNC)
| 3 | GND | HIF_HALL_SW | 5 | DC5V | 6 |  |  |  |  | FIL_RY_7M | 13 | +12V | 14 |  | 13 | FIL_RY_7M | 14 | - | IS_MOT | 2 |  |  |  |  | HIFU_RF | 1
|  |  | TXD_ACT | 7 | CABLE_SDA | 8 |  |  |  |  | FIL_RY_4M | 15 | +12V | 16 |  | 15 | FIL_RY_4M | 16 | - | [6 empty] | GND | 2
| CN105 |  | RXD_ACT | 9 | CABLE_SCL | 10 |  |  |  |  | RF_RY1 | 17 | FW_CT | 18 |  | 17 | RF_RY1 | 18 | FW_CT | CN205 |  | CN205
| 35313-2P |  | H12V | 11 | GND | 12 |  |  |  |  | RF_RY2 | 19 | RV_CT | 20 |  | 19 | RF_RY2 | 20 | RV_CT | 5268-8P |  | 5268-8P |  |  |  | RF_OUT2
| 1 | +12V | H12V | 13 | GND | 14 | [13 empty] | +12V | 1 | 1 | +12V |  |  | SMA-R-R/A(BNC)
| 2 | GND | [17 empty] | FIL_RY_7M | 2 | 2 | FIL_RY_7M |  |  | NED_RF_A | 1
|  |  | HD2 | [16 empty] | FIL_RY_4M | 3 | 3 | FIL_RY_4M |  |  | NED_RF_B | 2
| CN103 |  | NED_SDA | 1 | NS12V | 2 |  |  | CN401 |  | CN403 |  |  |  |  | CN201 |  |  |  | RF_RY1 | 4 | 4 | RF_RY1
| 5267-4P |  | NED_SCL | 3 | - | 4 |  |  | 5267-5P |  | MX5566-04P |  |  |  |  | MX5569-04P |  |  |  | RF_RY2 | 5 | 5 | RF_RY2
| 1 | +5V | NED_CART_SW | 5 | AGND | 6 |  |  | 1 | +5V | S12V | 1 | GND | 2 |  | 1 | +12V | 2 | GND | FW_CT | 6 | 6 | FW_CT
| 2 | TXD_DBG | NED_LIMIT_SW | 7 | AGND | 8 |  |  | 2 | BLUE_LED | S24V | 3 | GND | 4 |  | 3 | +24V | 4 | GND | RV_CT | 7 | 7 | RV_CT
| 3 | RXD_DBG | NED_SHOT_SW | 9 | NED_HP_SW | 10 |  |  | 3 | RED_LED | [9 empty] | GND | 8 | 8 | GND
| 4 | GND | MOT_1A | 11 | MOT_1B | 12 |  |  | 4 |  |  |  | CN405
|  |  | MOT_2A | 13 | MOT_2B | 14 |  |  | 5 | GND |  |  | 35313-2P
[13 empty] | +12V | 1
[5 empty] | CN104 |  |  |  | CN402 |  |  |  | GND | 2
[5 empty] | 35313-2P |  |  |  | LAB1143-2P
[5 empty] | +12V | 1 |  |  | 1 | S12V |  |  | CN404
[5 empty] | GND | 2 |  |  | 2 | GND |  |  | 35313-3P
[13 empty] | +24V | 1
[13 empty] | - | 2
[13 empty] | GND | 3
PCN
FRF
VDM211
(DC12V)
VDM211
(DC12V)
15RNS-L
(DC12V)
FLT
H/P
(HIFU)
H/P
(Needle)
PWR
CPU
H/P
(HIFU)
H/P
(Needle)
FOOT
S/W
FAN
(PWR)
FAN
(PWR)
LCD
10.2 inch
(ForStech)
FRONT LED
SMPS 12V
(MDS-100 BPS12)
SMPS 24V
(MDS-100 APB24)
LSW
DC/DC
24V to 12V
CT Block
Note on R13: 오타 +24V
Note on X24: 오타 +12V
Note on AH24: 오타 RF_OUT1

## Pin-map_HIFU_V0.1
Focus Dual Block Diagram
| CN101 |  | HD3 |  |  |  | HD1 |  |  |  | HD2 |  |  |  | HD2 |  |  |  | RF_OUT1
| 5267-2P |  | RF_QR_IN | 1 | PW_IS_CT | 2 | 1 | RF_QR_IN | 2 | PW_IS_CT | RF_FLT_OUT | 1 | RF_FLT_OUT | 2 | 1 | RF_FLT_OUT | 2 | RF_FLT_OUT | SMA-R-R/A(BNC)
| 1 | FOOT_SW | PW_PWM | 3 | VDC_OUT | 4 | 3 | PW_PWM | 4 | VDC_OUT | - | 3 | - | 4 | 3 | - | 4 | - | HIFU_BODY | 1
| 2 | GND | MOT_PWM | 5 | VDC_IN | 6 | 5 | MOT_PWM | 6 | VDC_IN | - | 5 | - | 6 | 5 | - | 6 | - | GND | 2
|  |  | PUMP_IS_CT | 7 | FW_CT | 8 | 7 | PUMP_IS_CT | 8 | FW_CT | GND | 7 | GND | 8 | 7 | GND | 8 | GND
| FAN1 |  | TIM1_OC_IN | 9 | RV_CT | 10 | 9 | - | 10 | RV_CT | GND | 9 | GND | 10 | 9 | GND | 10 | GND | RF_OUT2
| 5267-2P |  | TIM8_OC_IN | 11 | FET_TEMP | 12 | 11 | TIM8_OC_IN | 12 | FET_TEMP | - | 11 | - | 12 | 11 | - | 12 | - | SMA-R-R/A(BNC)
| 1 | +12V | BIAS_ON | 13 | PW_I_REF | 14 | 13 | BIAS_ON | 14 | PW_I_REF | - | 13 | - | 14 | 13 | - | 14 | - | HIFU_FACE | 1
| 2 | FAN_ON | VALVE2_ON | 15 | VALVE1_ON | 16 | 15 | VALVE2_ON | 16 | VALVE1_ON | SS24V | 15 | SS24V | 16 | 15 | SS24V | 16 | SS24V | GND | 2
|  |  | - | 17 | FILTER_7M | 18 | 17 | - | 18 |  | SS24V | 17 | SS24V | 18 | 17 | SS24V | 18 | SS24V
| FAN2 |  | HF_OUT1 | 19 | FILTER_4M | 20 | 19 | RELAY1 | 20 | FILTER_7M | - | 19 | - | 20 | 19 | - | 20 | - | CN202
| 5267-2P |  | HF_OUT2 | 21 | FILTER_2M | 22 | 21 | RELAY2 | 22 | FILTER_2M | [8 empty] | 53259-2P
| 1 | +12V | - | 23 | GND | 24 | 23 | RELAY3 | 24 | GND | HD3 |  |  |  | HD1 |  |  |  | +12V | 1
| 2 | FAN_ON | - | 25 | GND | 26 | 25 | - | 26 | GND | VAL1_COIL | 1 | VAL1_COIL | 2 | 1 | VAL1_COIL | 2 | VAL1_COIL | VAL1_COIL | 2
[11 empty] | VAL2_COIL | 3 | VAL2_COIL | 4 | 3 | VAL2_COIL | 4 | VAL2_COIL
| CN102 | [9 empty] | - | 5 | - | 6 | 5 | - | 6 | - | CN203
| 5267-3P |  | HD1 | [7 empty] | IS_MOT | 7 | IS_MOT | 8 | 7 | IS_MOT | 8 | IS_MOT | 53259-2P
| 1 | TXD_LCD | HIF_SHOT_SW | 1 | HIF_HP_SW | 2 |  |  |  |  | GND | 9 | GND | 10 | 9 | GND | 10 | GND | +12V | 1
| 2 | RXD_LCD | HIF_CART_SW | 3 | - | 4 |  |  |  |  | GND | 11 | GND | 12 | 11 | GND | 12 | GND | VAL2_COIL | 2
| 3 | GND | HIF_HALL_SW | 5 | DC5V | 6 |  |  |  |  | RF_RY3 | 13 | +12V | 14 | 13 | RF_RY3 | 14 | +12V
|  |  | TXD_ACT | 7 | CABLE_SDA | 8 |  |  |  |  | FIL_RY_4M | 15 | +12V | 16 | 15 | FIL_RY_4M | 16 | +12V | CN204
| CN105 |  | RXD_ACT | 9 | CABLE_SCL | 10 |  |  |  |  | RF_RY1 | 17 | FW_CT | 18 | 17 | RF_RY1 | 18 | FW_CT | LAB1143-2P
| 35313-2P |  | H12V | 11 | GND | 12 |  |  |  |  | RF_RY2 | 19 | RV_CT | 20 | 19 | RF_RY2 | 20 | RV_CT | +12V | 1
| 1 | +12V | H12V | 13 | GND | 14 | [12 empty] | IS_MOT | 2
| 2 | GND
|  |  | HD2 | [15 empty] | CN201
| CN103 |  | NED_SDA | 1 | NS12V | 2 | [12 empty] | LAB1143-3P
| 5267-4P |  | NED_SCL | 3 | - | 4 | [12 empty] | SS24V | 1
| 1 | +5V | NED_CART_SW | 5 | AGND | 6 | [12 empty] | - | 2
| 2 | TXD_DBG | NED_LIMIT_SW | 7 | AGND | 8 | [12 empty] | GND | 3
| 3 | RXD_DBG | NED_SHOT_SW | 9 | NED_HP_SW | 10
| 4 | GND | MOT_1A | 11 | MOT_1B | 12
|  |  | MOT_2A | 13 | MOT_2B | 14
CN401
| 5267-5P |  |  |  | CN104
| 1 | +5V |  |  | LAB1143-2P
| 2 | BLUE_LED |  |  | +12V | 1
| 3 | RED_LED |  |  | GND | 2
4
| 5 | GND
CN401
5267-5P
| 1 | +5V
| 2 | BLUE_LED
| 3 | RED_LED
4
| 5 | GND
VDM211
(DC12V)
VDM211
(DC12V)
15RNS-L
(DC12V)
H/P
(BODY)
H/P
(FACE)
PWR
CPU
H/P
(Body)
H/P
(Face)
FOOT
S/W
FAN
(PWR)
FAN
(PWR)
LCD
10.2 inch
(ForStech)
FRONT LED
SMPS 12V
(MDS-100 BPS12)
SMPS 24V
(MDS-100 APB24)
DC/DC
24V to 12V
CT Block
KEY LED
GPIO 부족한 경우, RELAY1만 사용(분기)
Water 용량 monitor : GPIO 2ea.
Water flow monitor : ADC 1ea.
SMA
DC 12V→5V
(LM2576S-5)
+5V
DC 12V→9V
(7809AF)
+9V
+12V

## RF&SMA_V0.1
|  | HD3 |  |  |  | HD1
|  | RF_QR_IN | 1 | PW_IS_CT | 2 | 1 | RF_QR_IN | 2 | PW_IS_CT
|  | PW_PWM | 3 | VDC_OUT | 4 | 3 | PW_PWM | 4 | VDC_OUT
|  | MOT_PWM | 5 | VDC_IN | 6 | 5 | MOT_PWM | 6 | VDC_IN
|  | PUMP_IS_CT | 7 | FW_CT | 8 | 7 | PUMP_IS_CT | 8 | FW_CT | [8 empty] | RF_OUT1
|  | TIM1_OC_IN | 9 | RV_CT | 10 | 9 | - | 10 | RV_CT | [8 empty] | SMA-R-R/A(BNC)
|  | TIM8_OC_IN | 11 | FET_TEMP | 12 | 11 | TIM8_OC_IN | 12 | FET_TEMP | [8 empty] | HIFU_BODY | 1
|  | BIAS_ON | 13 | PW_I_REF | 14 | 13 | BIAS_ON | 14 | PW_I_REF | [8 empty] | GND | 2
CN101 |  | W_PUMP_ON | 15 | SOL_ON | 16 | 15 | VALVE2_ON | 16 | VALVE1_ON
5267-2P |  | - | 17 | - | 18 | 17 | - | 18 | - | [8 empty] | RF_OUT2
1 | FOOT_SW | RELAY1 | 19 | FILTER_7M | 20 | 19 | RELAY1 | 20 | FILTER_7M | [8 empty] | SMA-R-R/A(BNC)
2 | GND | RELAY2 | 21 | - | 22 | 21 | RELAY2 | 22 | FILTER_2M | HD2 |  |  |  | HD2 |  |  |  | HIFU_FACE | 1
|  | RELAY3 | 23 | GND | 24 | 23 | RELAY3 | 24 | GND | RF_FLT_OUT | 1 | RF_FLT_OUT | 2 | 1 | RF_FLT_OUT | 2 | RF_FLT_OUT | GND | 2
|  | - | 25 | GND | 26 | 25 | - | 26 | GND | - | 3 | - | 4 | 3 | - | 4 | -
FAN1 | [9 empty] | - | 5 | - | 6 | 5 | - | 6 | - | CN201
5267-2P | [9 empty] | GND | 7 | GND | 8 | 7 | GND | 8 | GND | LAB1143-3P
1 | +12V | [8 empty] | GND | 9 | GND | 10 | 9 | GND | 10 | GND | SS24V | 1
2 | FAN_ON | [8 empty] | - | 11 | - | 12 | 11 | - | 12 | - | - | 2
[10 empty] | - | 13 | - | 14 | 13 | - | 14 | - | GND | 3
FAN2 | [9 empty] | SS24V | 15 | SS24V | 16 | 15 | SS24V | 16 | SS24V
5267-2P | [9 empty] | SS24V | 17 | SS24V | 18 | 17 | SS24V | 18 | SS24V
1 | +12V | [8 empty] | - | 19 | - | 20 | 19 | - | 20 | -
2 | FAN_ON
[10 empty] | HD3 |  |  |  | HD1 |  |  |  | CN202
CN103 | [9 empty] | VAL1_COIL | 1 | VAL1_COIL | 2 | 1 | VAL1_COIL | 2 | VAL1_COIL | 53259-2P
5267-4P | [9 empty] | VAL2_COIL | 3 | VAL2_COIL | 4 | 3 | VAL2_COIL | 4 | VAL2_COIL | +12V | 1
1 | +5V | [8 empty] | - | 5 | - | 6 | 5 | - | 6 | - | VAL1_COIL | 2
2 | TXD_DBG | [8 empty] | IS_MOT | 7 | IS_MOT | 8 | 7 | IS_MOT | 8 | IS_MOT
3 | RXD_DBG | [8 empty] | GND | 9 | GND | 10 | 9 | GND | 10 | GND | CN203
4 | GND | [8 empty] | GND | 11 | GND | 12 | 11 | GND | 12 | GND | 53259-2P
CN102 | [9 empty] | RF_RY3 | 13 | +12V | 14 | 13 | RF_RY3 | 14 | +12V | +12V | 1
5267-3P | [9 empty] | FIL_RY_4M | 15 | +12V | 16 | 15 | FIL_RY_4M | 16 | +12V | VAL2_COIL | 2
1 | TXD_LCD | [8 empty] | RF_RY1 | 17 | FW_CT | 18 | 17 | RF_RY1 | 18 | FW_CT
2 | RXD_LCD | [8 empty] | RF_RY2 | 19 | RV_CT | 20 | 19 | RF_RY2 | 20 | RV_CT | CN204
3 | GND | HD1 | [15 empty] | LAB1143-2P
|  | HP1_SHOT_SW | 1 | HP1_SW | 2 | [12 empty] | +12V | 1
CN105 |  | HP1_CART_SW | 3 | - | 4 | [12 empty] | IS_MOT | 2
35313-2P |  | HP1_HALL_SW | 5 | DC5V | 6
1 | +12V | HP1_TXD | 7 | HP1_SDA | 8
2 | GND | HP1_RXD | 9 | HP1_SCL | 10
|  | H12V | 11 | GND | 12
CN104 |  | H12V | 13 | GND | 14
LAB1143-2P
1 | +12V
2 | GND | HD2
|  | HP2_SHOT_SW | 1 | HP2_SW | 2
|  | HP2_CART_SW | 3 | - | 4
|  | HP2_HALL_SW | 5 | DC5V | 6
|  | HP2_TXD | 7 | HP2_SDA | 8
|  | HP2_RXD | 9 | HP2_SCL | 10
|  | H12V | 11 | GND | 12
|  | H12V | 13 | GND | 14
CN108
5267-2P
|  |  |  | WATER_LEVEL1 | 1
|  |  |  | GND | 2
CN106 |  |  |  | CN109
5267-4P |  |  |  | 5267-2P
1 | +5V |  |  | WATER_LEVEL2 | 1
2 | FRONT_BULE |  |  | GND | 2
3 | FRONT_RED
4 | GND |  |  | CN110
5267-2P
|  |  |  | WATER_TEMP | 1
CN107 |  |  |  | GND | 2
5267-5P
1 | +5V
2 | FRONT_KEY |  |  | CN111
3 | KEY_BULE |  |  | 35313-2P
4 | KEY_RED |  |  | +12V | 1
5 | GND |  |  | SOL_ON | 2
CN112
35313-2P
|  |  |  | +12V | 1
|  |  |  | W_PUMP | 2
buffer
DC 12V→5V
(LM2576S-5)
OPAmp
SPI
SPI
Amp
ULTRA-ENA
[3T:1T]
Transformer
FD-PWR-T4
MOSFET
IRFP450
MOSFET
IRFP450
[3T:1T]
Transformer
FD-PWR-T4
[3T:1T]
Transformer
FD-PWR-T4
DC24V
RF_OUT
DDS_CS
DDS_CLK
DDS_SDI
+12V
+5V
DC 12V→9V
(7809AF)
+9V
DC 24V→12V
(LM2576HV-12)
DC 12V→-12V
(TPS5430)
-12V
SS24V
S24V
+24V
MOSFET
IRF6218
LC Filter
DC24V
Driver
IR4427
PW_PWM
MOT_PWM
PWM_GATE
OPAmp
PW_IS_CT
VDC_IN
VDC_OUT
OPAmp
OPAmp
OPAmp
TIM8_OC_IN
PW_I_REF
Sig Gen
(AD9833)
Fuse
(10A/250VAC)
I-Sensor
(PP1S10-12)
Temp Sensor
OPAmp
Temp Sensor
OPAmp
FET_TEMP
VDM211
(DC12V)
VDM211
(DC12V)
15RNS-L
(DC12V)
BIAS_ON
Driver
(TD62783)
RF_RY1
RF_RY2
RF_RY3
FIL_RY_7M
RELAY1
RELAY2
RELAY3
FILTER_7M
Driver
IR4427
VALVE1_ON
VALVE2_ON
FET
FET
VAL1_COIL
VAL2_COIL
I-Sensor
(PP1S10-12)
OPAmp
PUMP_IS_CT
FET
IS_MOT
PWM_GATE
3MHz_SineWave
RF_INA
RF_AMP_OUT
RELAY
HL2-2CH-12
RF_RY1
RF_RY2
FIL_RY_7M
RELAY
HL2-2CH-12
FILTER
2MHz
FILTER
4MHz
FILTER
7MHz
RELAY
HL2-2CH-12
RF_FLT_OUT
CPU PCB에 배치 검토
PCN
CT
FT50-43
RF_FLT_OUT
FW_CT
RV_CT
RELAY
HL2-2CH-12
RF_RY3
+12V
+12V
+12V
+12V
HIFU_BODY
HIFU_FACE
FAN 1ea 추가
Body : 2MHz
Face : 4MHz, 7MHz
+12V
GPIO 부족한 경우, RELAY1만 사용(분기)
SMPS 24V
(MDS-100 APB24)
FOOT
S/W
FAN
(PWR)
FAN
(PWR)
LCD
10.2 inch
(ForStech)
FRONT LED
KEY LED
SMPS 12V
(MDS-100 BPS12)
TIM8_BKIN
PC5_POWER_ENA
PW_PWM
PWM_1A
MOT_ENA
MOT_PWM
DAC1
OPAmp
TP11
DAC2
OPAmp
PW_I_REF
RS-232
(MAX3232CSE)
UART1_TX
UART1_RX
UART3_TX
UART3_RX
TXD_LCD
RXD_LCD
TXD_DBG
RXD_DBG
I2C Buffer
3.3V -> 12V
(P82B96)
I2C2_SCL
I2C2_SDA
HP1_SCL
HP1_SDA
FW_CT_AD
1
FW_CT
RV_CT_AD
OPAmp
RV_CT
FET_TEMP_AD
OPAmp
FET_TEMP
PC4_PUMP_IS_AD
OPAmp
PUMP_IS_CT
VDC_IN_AD
OPAmp
VDC_IN
PW_IS_AD
OPAmp
PW_IS_CT
VDC_OUT_AD
OPAmp
VDC_OUT
PD2_HP1_SHOT_IN
HP1_SHOT_SW
PD1_HP1_CART_IN
HP1_CART_SW
Buffer
(74LS07)
+5V
+3.3V
PD0_HP1_HALL_IN
HP1_HALL_SW
PC12_HP1_IN
HP1_SW
PC11_FOOT_IN
FOOT_SW
Buffer
(74LS07)
+5V
+3.3V
PE15_TIM1_BKIN
TIM1_OC_IN
TIM8_BKIN
TIM8_OC_IN
PC8_BIAS_ON
BIAS_ON
PC7_W_PUMP_ON
W_PUMP_ON
Driver
(74HCT244)
+5V
+3.3V
PD15_FLT_7M
FILTER_7M
PD14_RLY1
RELAY1
PD13_RLY2
RELAY2
PE8_GPIO_ON
PD12_RLY3
PD11_HF_OUT2
HF_OUT2
PA11_NHP_PW_ON
NHP_PW_ON
Driver
(74HCT244)
+5V
+3.3V
PD10_EPP_ON
EPP_ON
PE11_HP_RY_ON
HP_RY_ON
PC10_SOL_ON
SOL_ON
PE8_GPIO_ON
PE10_FAN_ON
FAN_ON
FAN_IS_AD
1
FAN_ON
DC 12V→5V
(NCP3064)
+12V
+5V
DC 5V→3.3V
(7809AF)
+3.3V
DC 12V→5V
(78L05F)
DC5V
EPP_ON
BJT
RELAY
HL2-2CH-12
H12V
HP_RY_ON
+12V
RELAY3
I2C Buffer
3.3V -> 12V
(P82B96)
I2C1_SCL
I2C1_SDA
HP2_SCL
HP2_SDA
PD7_HP2_SHOT_IN
HP2_SHOT_SW
PD6_HP2_CART_IN
HP2_CART_SW
Buffer
(74LS07)
+5V
+3.3V
PD5_HP2_HALL_IN
HP2_HALL_SW
PD4_HP2_IN
HP2_SW
H/P
(BODY)
H/P
(FACE)
WATER_LV1_AD
OPAmp
WATER_LEVEL1
WATER_LV2_AD
OPAmp
WATER_LEVEL2
WATER_TEMP_AD
OPAmp
WATER_TEMP
SOL Valve
( )
Water Pump
( )
Water Temp Sensor
( )
Water Level Sensor
( )
Water Level Sensor
( )
EEPROM
(24C02)
I2C3_SCL
I2C3_SDA
KEY_RED
KEY_BULE
FRONT_RED
FRONT_BULE

## HP
HIFU OUT+
| HIFU OUT- | [5 empty] | 엑츄에이터 연결
CN604
| HD1 |  |  |  | CN601 |  | 1 | TXD_SER
| HIF_SHOT_SW | 1 | HIF_HP_SW | 2 | 1 | +12V | 2 | RXD_SER
| HIF_CART_SW | 3 | - | 4 | 2 | +12V | 3 | +7V
| HIF_HALL_SW | 5 | DC5V | 6 | 3 | DC5V | 4 | GND
| TXD_ACT | 7 | CABLE_SDA | 8 | 4 | GND
| RXD_ACT | 9 | CABLE_SCL | 10 | 5 | GND | LED PCB 연결
| H12V | 11 | GND | 12 |  |  | HD2
| H12V | 13 | GND | 14 | CN602 |  | 1 | RED1 | 2 | +5V
[5 empty] | 1 | TXD_ACT | 3 | RED2 | 4 | -
1 | HIF_SHOT_SW | Shot S/W input |  |  | 2 | RXD_ACT | 5 | BLUE1 | 6 | GND
2 | HIF_CART_SW | Cart indicator |  |  | 3 | CABLE_SDA | 7 | BLUE2 | 8 | GND
3 | HIF_HP_SW | HP Cable 연결 detact |  |  | 4 | CABLE_SCL
4 | CABLE_SDA | EEPROM/IO Expandor 통신 |  |  |  |  | Kartridge
5 | CABLE_SCL | EEPROM/IO Expandor 통신 |  |  | CN603 |  | HD1
6 | H12V |  |  |  | 1 | SHOT_SW | 1 | EP_SCL | 2 | DC5V
7 | DC5V |  |  |  | 2 | KART_CHK | 3 | EP_SDA | 4 | +5V
8 | GND |  |  |  | 3 | HALL_CHK | 5 | KART_CHK | 6 | GND
9 | HIFU OUT+ |  |  |  | 4 | GND | 7 | SEN_POS | 8 | GND
10 | HIFU OUT-
5 | HIF_HALL_SW | Cart actuator 위치 check
7 | TXD_ACT | Actuator 통신
9 | RXD_ACT | Actuator 통신
13 | H12V
14 | GND
4 | -
I2C Buffer
(P82B96)
TXD_SER
RXD_SER
EP_SCL
EP_SDA
CABLE_SCL
CABLE_SDA
TXD_ACT
RXD_ACT
+12V
HALL_CHK
BJT
I/O Expandor
(PCF8574AT)
2-wire serial bus extender
(PCA9507D)
+5V
+5V
LED S/W
SEN_POS
I2C Buffer
(P82B96)
+12V
DC/DC
(78L05F)
+12V
+5V
DC/DC
(LD29150PT)
+7V

## Check
Review Data
전원 Key type
Push 버튼인 경우, LSW PCB 필요함.
Key S/W인 경우, LSW PCB 필요없음.
전원 분리 : HIFU 장비는 인체에 직접 접촉되지 않기 때문에 필요 없음.
LED PCB 필요 부분
Foot S/W 지원 여부